## Title
Medical Physics and Clinical Engineering |  |  |  | Addenbrookes Hospital
Clinical Engineering / R&D
Risk File |  |  |  | CE-ER-FRM-22
|  | Product Identification | Surgical Mask Clips
Document Q-Pulse reference
|  | Doc. Name | Risk File - Mask Clips
|  | Mandatory Supporting Documents | Q-Pulse reference
Risk Management Process
Risk Plan
Risk Report
Product description and intended use
The product is a disposable attachment which is used to secure the elastic ear loops of disposable surgical masks by providing an attachment point on the back of the head. It is intended to be used: To improve security of disposable Type IIR surgical masks with elastic ear loops against face compared to having no extra securement method To be more comfortable than the current recommended securement method (knots tied in the ear loops) Type IIR surgical masks are a Class 1 medical device, and as such the face mask clip is an accessory to a medical device.
Cambridge University Hospitals NHS Foundation Trust |  |  |  |  | Page 1 of 4
This document complies with the Cambridge University Hospitals NHS Foundation Trust service Equality and Diversity statement and must be read in conjunction with Addenbrookes Hospital Policies & Procedures. It is your responsibility to check against the electronic library that this printed out copy is the most recent issue of this document. All changes required should be submitted via Q-Pulse Document module (Change request).
Filename: 008_Risk Evaluation Table
Document No.: CE-ER-FRM-15 |  |  | Version No.: 5 | Active on Q-Pulse & Published

## Revision Table
Document History
Date | Identification | Description | Name
2020-08-04 | 1.0 | Created (Draft) | Sonya Sireau

## Risk Evaluation Criteria
CUH Risk management strategy and policy 2019-2019
CUH Risk management strategy and policy 2019-2020
Source | CUH Risk management strategy and policy 2019-2020, http://merlin/Lists/DMSRecords/DispRecordTabsDoc.aspx?ID=18427
Adaptations | The risk scoring matrix and grading system has been adapted to enable the CEI team to better track and manage risks that are yet to be determined, or cannot be determined.
Notes | Risk evaluation is for injury or damage to the health of people, or damage to property or the environment, and does include risks surrounding data security. It should NOT consider project or financial impacts. These should instead be considered in the Project Viability Assessment.
Risk Evaluation
| Severity Score |  | Likelihood Score | [6 empty] | Grading Risk
|  |  | 1 | 2 | 3 | 4 | 5 | TBD | CBD | 1-3 | Low risk | No action required
|  |  | Rare | Unlikely | Possible | Likely | Almost Certain | To be determined | Cannot be determined | 4-6 | Moderate risk | Risk reduction if practicable
| 1 | Negligible | 1 | 2 | 3 | 4 | 5 | 25 | 25 | 8-12 | High risk | Risk reduction if practicable, cost benefit analysis required for risks that cannot be practibably lowered
| 2 | Minor | 2 | 4 | 6 | 8 | 10 | 25 | 25 | 15-25 | Significant risk | Unnacceptable risk
| 3 | Moderate | 3 | 6 | 9 | 12 | 15 | 25 | 25 | TBD | To be determined | Activities required to define risk likelihood or severity
| 4 | Major | 4 | 8 | 12 | 16 | 20 | 25 | 25 | CBD | Can't be determined | Treat as unnacceptable and define control measures, or justify lower risk score
| 5 | Catastrophic | 5 | 10 | 15 | 20 | 25 | 25 | 25
| TBD | To be determined | 25 | 25 | 25 | 25 | 25 | 25 | 25
| CBD | Cannot be determined | 25 | 25 | 25 | 25 | 25 | 25 | 25
Likelihood Score
|  | 1 | 2 | 3 | 4 | 5
| Descriptor | Rare | Unlikely | Possible | Likely | Almost Certain
| Frequency | This will probably never happen/recur | Do not expect it to happen/recur but it is possible | Might happen or recur occasionally | Will probably happen/recur but it is not a persisting issue | Will undoubtedly happen/recur,possibly frequently
Severity Score
|  | 1 | 2 | 3 | 4 | 5
| Descriptor | Negligible | Minor | Moderate | Major | Catastrophic
| Impact on the safety of patients, staff or public (physical/psychological harm) | Minimal injury requiring no/minimal intervention or treatment. | Minor injury or illness, requiring minor intervention. Increase in length of hospital stay by 1-3 days Requiring time off work for >3 days | Moderate injury requiring professional intervention. Requiring time off work for 4-14 days Increase in length of hospital stay by 4-15 days RIDDOR/agency reportable incident An event which impacts on a small number of patients. | Major injury leading to long-term incapacity/disability. Increase in length of hospital stay by >15 days. Mismanagement of patient care with longterm effects. Requiring time off work for >14 days | Incident leading to death. Multiple permanent injuries or irreversible health effects. An event which impacts on a large number of patients.
| Quality of treatment | Peripheral element of treatment or service suboptimal Informal complaint/inquiry | Overall treatment or service suboptimal Single failure to meet internal standards | Treatment or service has significantly reduced in effectiveness Repeated failure to meet internal standards | Non-compliance with national standards with significant risk to patients if unresolved Multiple complaints/ independent review Low performance rating Critical report | Totally unacceptable level or quality of treatment/service Gross failure of patient safety if findings not acted on Inquest/ombudsman inquiry Gross failure to meet national standards
| Personal information | No risk to confidential staff or patient data | Occassional/one-off cases of personal data exposure within CUH | Data exposure within CUH | Public exposure of personal data Exposure of personal data to entities who may use that data against the best interest of the person to whom the data belongs | Large scale public exposure of personal data Exposure of personal data to entities who will use that data against the best interest of the person to whom the data belongs
| Envrionmental impact | Minimal or no impact on the environment | Minor impact on environment | Moderate impact on environment | Major impact on environment | Catastrophic impact on environment
| Environmental impact | Minimal or no impact on the environment | Minor impact on environment | Moderate impact on environment | Major impact on environment | Catastrophic impact on environment

## Benefit Evaluation Criteria
008_Risk Evaluation Table
Benefit Evaluation Criteria
Source | Developed by CEI
Adaptations | n/a
Notes | Benefit evaluation is for assessing impact on positive impact on clinical outcome, the patient’s quality of life, outcomes related to diagnosis, positive impact from diagnostic devices on clinical outcomes, or positive impact on public health or patient management. It should NOT consider project or financial impacts. These should instead be considered in the Project Viability Assessment.
Risk Evaluation
| Benefit Size Score |  | Certainty Score |  |  |  |  | Grading Risk
|  |  | 1 | 2 | 3 | 4 | 5 | 15-25 | Large, clear benefit
|  |  | Very unlikely to be realised | Unlikely | Possible | Likely | Almost Certain | 2020-10-12 | Reasonable benefit
| 1 | Negligible | 1 | 2 | 3 | 4 | 5 | 2020-05-09 | Small benefit
| 2 | Minor | 2 | 4 | 6 | 8 | 10 | 2020-01-04 | Very little benefit
| 3 | Moderate | 3 | 6 | 9 | 12 | 15
| 4 | High | 4 | 8 | 12 | 16 | 20
| 5 | Very High | 5 | 10 | 15 | 20 | 25
Likelihood Score
|  | 1 | 2 | 3 | 4 | 5
| Descriptor | Rare | Unlikely | Possible | Likely | Almost Certain
| Frequency | Possible but quite unlikely outcome of the project/extremely uncertain outcome of the project | Possible outcome of the project, but uncertain | Possible outcome of the project | A likely outcome of the project | Will undoubtedly happen as an outcome of the project if completed successfully
Impact Score
|  | 1 | 2 | 3 | 4 | 5
| Descriptor | Negligible | Minor | Moderate | High | Very High
| Medical Benefit | No measurable impact on medical outcomes No measurable impact on patient experience Minor anecdotal benefit | Minor measurable impact on medical outcomes Minor measurable impact on patient experience Moderate anecdotal benefit | Moderate measurable impact on medical outcomes Moderate measurable impact on patient experience Large anecdotal benefit (e.g. patients happy to engage with a home treatment method where they otherwise wouldn't) | Significant measurable impact on key medical outcomes Significant measurable impact on patient experience | Measurable life-saving impact
| Scale of impact | Hospital, short term or one-off Less than 100 patients per year | Hospital Hundreds of patients per year | Hospital or regional Thousands of patients per year | National 10s of thousands of patients per year | International 100s of thousands of patients affected per year
| Clinical benefit | No measurable impact | Minor increase in number of patients treated for a given resource (e.g. reduced hospital stay Minor decrease in resource required to treat a given number of patients (lower cost per intervention, reduction in consultant time per appointment) Minor improvement in service quality (e.g. speed of diagnosis, error reduction) | Moderate increase in number of patients treated for a given resource (e.g. reduced hospital stay Moderate decrease in resource required to treat a given number of patients (lower cost per intervention, reduction in consultant time per appointment) Moderate improvement in service quality (e.g. speed of diagnosis, error reduction) | Large increase in number of patients treated for a given resource (e.g. reduced hospital stay Large decrease in resource required to treat a given number of patients (lower cost per intervention, reduction in consultant time per appointment) Large improvement in service quality (e.g. speed of diagnosis, error reduction) | Transformative increase in number of patients treated for a given resource (e.g. reduced hospital stay Transformative decrease in resource required to treat a given number of patients (lower cost per intervention, reduction in consultant time per appointment) Transformative improvement in service quality (e.g. speed of diagnosis, error reduction)

## Overall Benefit Risk
008_Risk Evaluation Table
Benefit Risk Summary
Surgical Mask Clip
| Overall Benefit: | 18
| Residual Risk: | 5
| Summary of benefit: | Surgical masks with ear loops need to be secured more firmly to the face than ear loops allow - greater security of the mask has a moderate to significant impact on the risk of transmission. This solution enables this security to take place without causing additional discomfort to the wearer compared to wearing an un-secure mask.
| Summary of risk: | There are two areas of risk which may increase the risk of COVID-19 transmission to staff compared with continuing to use knots. First is the residual risk that COVID-19 may be transmitted from manufacturer to user via the clip surface. Second is the residual risk associated with increased handling of hair during doffing if the clip becomes tangled.
| Actions summary: | Continue to evaluate designs whilst in use.
Optional: Benefit-risk summary for current state-of-the-art
Knotted ear loops
| Overall Benefit: | 16
| Residual Risk: | 20
| Summary of benefit: | Surgical masks with ear loops need to be secured more firmly to the face than ear loops allow - greater security of the mask has a moderate impact on the risk of transmission. This solution enables the mask to be secured to the face much better than not having knots in the ear loops.
| Summary of risk: | Discomfort from the tightened ear loops is causing pain/injury amongst a large proportion of staff. For some staff, this is increasing level of strees and affecting ability to provide care. For a minority of staff, there is a refusal to accept these masks and staff are insisting on FFP3 masks instead. This could cause future shortages through unnecescary use of PPE. It is also likely that some staff will bring in thei own alternatives which carry additional risk
Actions summary:
NOT APPLICABLE - product proposal is due to absence of availability of state-of-the-art

## Device Benefit Evaluation
008_Risk Evaluation Table
Benefit Evaluation Table for medical and clinical benefits associated with the medical device
Source | Developed by CEI
| Overall Benefit: | 18 | Evaluate overall residual medical/clinical benefit. This can be estimated as the largest single RPN score.
S = severity, L = likelihood, RPN = Risk Priority Number
Unique ID for hazard plus control measure - useful for referencing hazards in other documents | Brief note explaining benefit | Brief note explaining consequence of benefit | Define any investigatory activities/follow up actions when it is too early to define impact/likelihood | Impact | Likelihood | Score
ID | Benefit | Consequence | Investigatory activities required | S | L | RPN
FV01 | Surgical masks are held more securely to face, increasing efficacy and decreasing risk of transmission of virus from staff to patient/patinet to staff. | Moderate reduction of staff risk in contracting COVID-19 |  | 4 | 4 | 16
FV02 | Masks are secured more comfortably than alternative methods. | Greater adherence to protocol, reduction in risk of staff contracting covid. |  | 4 | 3 | 12
FV03 | Masks are secured more comfortably than alternative methods. | Increased comfort |  | 2 | 5 | 10
FV04 | Masks are secured more comfortably than alternative methods. | Staff will not insist on wearing FFP3 masks to prevent ear pain, so less PPE will be wasted. This may prevent future infections casued by lack of PPE, or inadequate care due to lack of PPE. |  | 4 | 3 | 12
0
0
0
Benefit Evaluation Table for medical and clinical benefits associated with best alternative (knots in ear loops)
Source | Developed by CEI
| Overall Benefit: | 16 | Evaluate overall residual medical/clinical benefit. This can be estimated as the largest single RPN score.
S = severity, L = likelihood, RPN = Risk Priority Number
Unique ID for hazard plus control measure - useful for referencing hazards in other documents | Brief note explaining benefit | Brief note explaining consequence of benefit | Define any investigatory activities/follow up actions when it is too early to define impact/likelihood | Impact | Likelihood | Score
ID | Benefit | Consequence | Investigatory activities required | S | L | RPN
FV01 | Surgical masks are held more securely to face, increasing efficacy and decreasing risk of transmission of virus from staff to patient/patinet to staff. | Moderate reduction of staff risk in contracting COVID-19 |  | 4 | 4 | 16
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0
0

## Device Risk Evaluation - clips
Risk Evaluation Table for health hazards associated with the medical device that may impact users or patients
Scope: | injury or damage to the health of people, or damage to property or the environment
Source | Developed by CEI
| Device: | Surgical Face Mask Clips
| Residual Risk: | 5 | Evaluate overall residual medical/clinical risk following the process described in the Risk Management Process CE-ER-PRO-14
Explanation:
S = severity, L=likelihood
Unique ID for hazard plus control measure | Brief note describing hazard | Brief note expalining impact/consequences of hazard | Severity | Likelihood | Score | Define any investigatory activities/follow up actions when it is too early to define a risk control activity | Define activities to control risk (or summarise existing evidence) | Provide a Q-Pulse reference, or location reference to the Technical File | Provide a date of completion for risk control activitie, to confim the activity has been completed | Severity | Likelihood | Score | Detail outcome - is the risk acceptable? If not, is a further risk control measure required (please enter the same hazard with a different ID and repeat the risk evaluation and control prcoess)? Alternatively, is a risk benefit analysis required? | Optional - tick if a risk/benefit is required | Please summarise the outcome of the risk/benefit analysis | If not in this document, please supply a Q-Pulse reference or location reference to the Technicla Filefor location of risk-benefit analysis
ID | Hazard | Consequence | S | L | RPN | Further activities required | Risk Control Measures | Supporting Documentation | Date carried out dd/mm/yyyy | S2 | L2 | RPN2 | Outcome | Risk/Benefit required? | Risk Benefit Outcome | Supporting Documentation2
FV01 | Interference with donning of PPE - Excessive touching of hair, or increased likelihood of contacting inside of mask compared to no clip. | Increase likelihood of clinician contacting COVID. Note, they should be donning PPE with cleaned, non-gloved hands. | 5 | 1 | 5 |  | Design & user testing | Verification & Validation | 20/04/2020 | 5 | 1 | 5 | Unacceptable | Yes | Risk acceptable
FV02 | Interference with doffing of PPE - increased contact time with hair, increased likelihood of contacting outside surface of mask. Note, PPE will be doffed using clean, non-gloved hands. | Increase likelihood of clinician contracting COVID. | 5 | 2 | 10 |  | Design & user testing. Instructions for use showing pictures of donning and doffing | Verification & Validation | 20/04/2020 | 5 | 1 | 5 | Unacceptable | Yes | Risk acceptable
FV03 | Interference with doffing of PPE - clip falls off onto floor before being placed in bin. Note, PPE will be doffed using clean, non-gloved hands. | Increase likelihood of clinician contacting COVID, since clinician will have to handle unclean clip to transfer to bin. | 5 | 2 | 10 | Is there a protocol in place for e.g. surgical masks that miss the bin | Design & user testing | Verification & Validation | 20/04/2020 | 5 | 1 | 5 | Unacceptable | Yes | Risk acceptable
FV04 | User hair snagging | Hands are clean, but hair may provide a transmission surface to clean hands. Worst case - contract COVID, death. | 5 | 2 | 10 |  | Design selection & user testing | Verification & Validation | 20/04/2020 | 5 | 1 | 5 | Unacceptable | Yes | Risk acceptable
FV05 | Catches with cap | Increased handling of dirty PPE when doffing. | 5 | 2 | 10 |  | Design & user testing | Verification & Validation | 20/04/2020 | 5 | 1 | 5 | Unacceptable | Yes | Risk acceptable
FV06 | Interfers with visor / cap | Decreases effectiveness of other PPE. | 5 | 1 | 5 |  | Design & user testing | Verification & Validation | 20/04/2020 | 5 | 1 | 5 | Unacceptable | Yes | Risk acceptable
FV07 | Clip breaks | Foreign object in clinical environment | 3 | 3 | 9 | Is there an existing protocol for this? | Material properties Geometry designed to avoid stress concentration Designed to handle sweat QC to avoid manufacturing defects that may lead to breakage | Verification & Validation | 20/04/2020 | 3 | 1 | 3 | Acceptable | Yes | Risk acceptable
FV08 | Viral transmission from manufacturer to user via persistence of virus on device materials | Infection of clinical staff member, who may infect other staff or patients - worst possible consequence is death | 5 | 3 | 15 | Check longevity of viable virus in literature | Process control to ensure risk of transmission from manufacturer to product is minimised | Manufacturing SOP | 22/04/2020 | 5 | 1 | 5 | Unacceptable | Yes | Risk acceptable
FV09 | Viral transmission due to presence of virus on material for visor before arrival at the manufacturing facility. Transmission to manufacturing staff, or to clinical staff. | Infection of clinical staff member, who may infect other staff or patients - worst possible consequence is death | 5 | 3 | 15 | Risk assess different material supply options | Prevention through material handling controls | Manufacturing SOP | 22/04/2020 | 5 | 1 | 5 | Unacceptable | Yes | Risk acceptable
FV10 | Material of mask clip irritates skin | Skin damage | 2 | 3 | 6 |  | Ensure material selected is appropriate for skin contact up to 4 hr | Waiting on data sheets - however, low risk that food grade cardboard and food grade polypropylene will cause harm | 22/04/2020 | 2 | 2 | 4 | Acceptable | Yes | Risk acceptable
FV11 | Mask does not fit face as well as without the clip | Increase likelihood of clinician contracting COVID. | 5 | 2 | 10 | Initial evaluation by range of staff required - could we put a QR code on visor to a GoogleForm? | Testing with a range of different head size users . Clips are designed so that masks fit better | Verification & Validation | 20/04/2020 | 5 | 1 | 5 | Unacceptable | Yes | Risk acceptable
FV12 | Clip intereferes with hairstyle (e.g. ponytail) and fit is poor on face | Increase likelihood of clinician contracting COVID. | 5 | 3 | 15 | Initial evaluation by range of staff required - could we put a QR code on visor to a GoogleForm? | Testing with a range of different hairstyles. Clips can be worn in a variety of ways to accomodate different hairstyles | Verification & Validation | 20/04/2020 | 5 | 1 | 5 | Unacceptable | Yes | Risk acceptable
FV13 | Sharp edges cause cut when being handled | Delay in clincian seeing patient whilst cut is cleaned and covered. | 3 | 2 | 6 |  | Quality control during manufacturing process Design | Verification & Validation | 20/04/2020 | 3 | 1 | 3 | Acceptable
FV14 | Clinician handles mutiple clips when picking own clip up | Clinican to clinician transmission of COVID via clips wiating to be used. Worst scenario: death. | 5 | 2 | 10 | Discussion with infection control | Standard practice is that clinicians will all have clean ahnds when donning ppe | Device prescrition (sign off of design for container from Infection Control) | 22/04/2020 | 5 | 1 | 5 | Unacceptable | Yes | Risk acceptable
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0

## Device Risk Evaluation - Knots
Risk Evaluation Table for health hazards associated with the medical device that may impact users or patients
Scope: | injury or damage to the health of people, or damage to property or the environment
Source | Developed by CEI
| Device: | Knotted ear loops
| Residual Risk: | 15 | Evaluate overall residual medical/clinical risk following the process described in the Risk Management Process CE-ER-PRO-14
Explanation:
S = severity, L=likelihood
Unique ID for hazard plus control measure | Brief note describing hazard | Brief note expalining impact/consequences of hazard | Severity | Likelihood | Score | Define any investigatory activities/follow up actions when it is too early to define a risk control activity | Define activities to control risk (or summarise existing evidence) | Provide a Q-Pulse reference, or location reference to the Technical File | Provide a date of completion for risk control activitie, to confim the activity has been completed | Severity | Likelihood | Score | Detail outcome - is the risk acceptable? If not, is a further risk control measure required (please enter the same hazard with a different ID and repeat the risk evaluation and control prcoess)? Alternatively, is a risk benefit analysis required? | Optional - tick if a risk/benefit is required | Please summarise the outcome of the risk/benefit analysis | If not in this document, please supply a Q-Pulse reference or location reference to the Technicla Filefor location of risk-benefit analysis
ID | Hazard | Consequence | S | L | RPN | Further activities required | Risk Control Measures | Supporting Documentation | Date carried out dd/mm/yyyy | S2 | L2 | RPN2 | Outcome | Risk/Benefit required? | Risk Benefit Outcome | Supporting Documentation2
FV01 | Pain caused by tightened ear loops | Discomfort/injury | 2 | 4 | 8
FV02 | Pain caused by tightened ear loops | Increased stress, impacting ability to do job | 3 | 4 | 12
FV03 | Pain caused by tightened ear loops | Refusal to wear surgical mask, only FFP3 masks. This may cause PPE shortages, resulting in inadequate protection for staff , and increased risk of infection. | 5 | 3 | 15
FV04 | Pain caused by tightened ear loops | Refusal to wear surgical mask, only FFP3 masks. This may cause PPE shortages, resulting in staff refusal to enter wards and inadequate care. | 5 | 3 | 15
FV05 | Process of knotting | Increased handling of mask while donning, increased likelihood of touching inside of mask and contracting COVID. | 5 | 1 | 5
FV06 | Staff bring in own ear protectors made without any safety controls in place | Ear protectors are reused or enter the trust infected with COVID. Some open source designs are not designed to facilitate safe donning / doffing. | 5 | 4 | 20
0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0
[5 empty] | 0 | [6 empty] | 0

## PHA
| Risk Questions | N/A | Detail
1.2 | Is the PPE intended to contact the patient or other persons? | Yes | Contact for up to 4 hours with hair and skin around forehead
1.3.1 | Are all materials used in the PPE biocompatible? | TBD | Materials have not been selected for biocompatibility, but are considerd to be low risk
1.3.2 | Will any harmful materials or substances be in contact with the patient? | No
1.4 | Is energy delivered to and/or extracted from the patient? | No
1.5 | Are substances delivered to and/ or extracted from the patient? | No
1.6 | Are biological materials processed by the PPE for subsequent re-use? | No
1.7.1 | Is the PPE supplied sterile or intended to be sterilised by the user? | No
1.7.2 | Are other microbiological controls applicable? | Yes | Meaures are taken through manufacturing process to minimise risk of virus being present on the clip, but cannot be 100% guaranteed
1.8 | Is the PPE intended to be routinely cleaned and disinfected by the user? | No
1.9 | Is the PPE intended to modify the patient environment? | No
1.1 | Are measurements taken by the device? | N/A
1.11 | Is the PPE interpretative? | N/A
1.12.1 | Is the PPE intended for use in conjunction with medicines? | N/A
1.12.2 | Is the PPE intended for use in conjunction with other medical technologies? | N/A
1.13 | Is there any unwanted output of energy or substances? | No
1.14 | Is the PPE susceptible to environmental influences? | No
1.15 | Does the PPE influence the environment? | No
1.16 | Are there essential consumables or accessories associated with the PPE? | No
1.17.1 | Is maintenance necessary? | No | To be single use or disposed of in event of failure to perform adequately.
1.17.2 | Is calibration necessary? | No
1.18 | Does the PPE contain software? | No
1.19 | Does the PPE have a restricted shelf-life? | No
1.2 | Are there any delayed and/or long-term use effects? | No
1.21 | Will the PPE be subject to mechanical forces? | No | No forces beyond routine operation - putting on and taking off.
1.22.1 | Does the PPE have a battery that will affect the lifetime of the device? | No
1.22.2 | Will ageing determine the lifetime of the PPE? | No
1.23 | Is the PPE intended for single use? | Yes
1.24 | Is safe decommissioning or disposal of the PPE necessary? | Yes | Clinical waste after use.
1.25.1 | Does installation of the PPE require special training? | N/A
1.25.2 | Does use of the PPE require special training? | No | Intuitive to use.
1.26 | Will new manufacturing processes need to be established or introduced? | Yes | See SOP created for manufacturing
1.27 | Is successful application of the PPE critically dependant on human factors? | No
1.27.1 | Does the PPE have connecting parts or accessories? | No
1.27.2 | Does the PPE have a control interface? | No
1.27.3 | Does the PPE display information? | No
1.27.4 | Is the PPE controlled by a menu? | No
1.28 | Is the PPE intended to be mobile or portable? | Yes | By design

## Classification
Accessory to Class I Medical Device.